## Slide 1: Calibration and Maintenance of Medical Device To Ensure Patient Safety
Dr. Biraj Chandra Paul
Assistant Professor, Hospital Administration
AIIMS, Guwahati

## Slide 2: PRIMUM NON NOCERE

## Slide 3: Introduction:
Modern day hospital with the advancement of technology, provides an advantage in that, it enhances the speed and accuracy of diagnosis, and efficacy of therapeutic procedures.
These benefits can be availed by patients only, if the medical devices are operationally reliable, is properly calibrated for precise and accurate functioning as well as free from hazards.
Instances of babies being roasted alive due to incubator burns, or hip implants causing blood poisoning, or burns due to faulty functioning of electric cautery, or short wave diathermy, or death of a patient due to malfunctioning ventilator are not unknown.
The medical devices if not properly functioning or not properly operated, can do a lot of damage and become a serious safety hazard for the patients.
It is therefore responsibility of the healthcare providers to ensure fault free functioning and full operational reliability of the medical devices in use

## Slide 4: The Medical device regulation in India
The Central Drugs Standard Control Organization (CDSCO), operating under the Ministry of Health and Family Welfare, serves as the primary regulatory authority for medical devices in India, establishing and enforcing standards that govern their manufacture, import, sale, and distribution.
A significant milestone was reached in 2017 with the introduction of the Medical Devices Rules, which provided a dedicated regulatory framework for medical devices. The rules establish four risk classes (A through D) with corresponding levels of regulatory control. Licensing procedures for both domestic manufacturers and importers, including the documentation requirements and timelines. In April 2020, these rules were amended to extend regulatory control to all medical devices, expanding beyond the 37 categories that were previously regulated.
The Drugs and Cosmetics Act, 1940, continues to serve as the primary legislation governing medical devices in India. The Act was amended multiple times to accommodate the evolving nature of medical devices. Section 3(b)(iv) of the Act specifically includes medical devices within its scope, defining them as instruments, apparatus, appliances, or materials intended for use in the diagnosis, treatment, mitigation, or prevention of diseases or disorders in human beings. It establishes penalties for violations and empowers regulatory authorities to take action against non-compliance.

## Slide 5: The Medical device regulation in India
The New Drugs and Clinical Trials Rules, 2019, introduced additional requirements for clinical investigations of medical devices.
In 2020, the government implemented a comprehensive scheme for voluntary registration of medical devices, which later became mandatory for specific categories of devices.
Compliance with ISO 13485 is mandatory for manufacturers of Class C and D devices, and strongly recommended for other classes.

## Slide 6: Materiovigilance program of India

## Slide 7: Measures for Operational Safety of Medical Devices:
There is a documented program of equipment management
Equipment management committee
Equipment audit committee
Efficient Engineering Service Department with biomedical equipment section with workshop facility for repair and maintenance of electrical, mechanical, electronic and biomedical equipment. Adequate trained staff and spares.
Mandatory reporting of accidents related to medical devices
Due care is given during installation such as temperature, humidity, voltage level recommendation.
Training of staff in using
Preventive and breakdown maintenance
Documented schedule of maintenance (daily, monthly, quarterly, and annual)
Procedure and schedule for recalibration of equipment requiring periodic recalibration
Following regulatory guidelines/ standards (IEC 60601, IEC 62353/ IS 8607, NABL 126

## Slide 8: Measures for Operational Safety of Medical Devices:
Regular energy audit
Regular equipment audit
Regular inspection of all biomedical equipment in patient care areas
Regular check of all electric cables, switches, earthing.
No lose/temporary electric wires
PPE is made available for use by the staff
Any equipment found operationally not reliable is declared redundant and removed from the usage area until its full function and operational reliability is restored.
Any known hazard of any equipment are being informed to all the user staff and warning sign in local language are displayed prominently

## Slide 9: National Biomedical Equipment Management & Maintenance (BMMP)
The MOHFW, GoI launched BMMP in the year 2015 for comprehensive maintenance of medical equipment available in the public health facilities from the level of PHC up to the district hospital
The program supports States/ UTS both in-house and PPP mode to maintain equipment inventory, repair and undertake CMC/ AMC for keeping maximum equipment functional in the health facility.
A good equipment management program demands a multidisciplinary equipment management committee at the core of the program which can provide a participatory mechanism for prudent, need based, purchases, planned maintenance with minimum downtime and optimal utilization with maximum returns
The in-house mode of service delivery is based on maintenance through OEMs by engaging through CMC or AMC. Financial support is provided under NHM through Annual Programme Implementation Plans (PIPs).
KPI for BMMP Equipment CAMC/ AMC (in percentage) and Equipment Up-Time (in percentage)

## Slide 10: Hospital Equipment Management Committee (HEMC)
Suggested Composition:
MS/ AMS or MO IC health facility
Head or Nominated Members from Clinical Specialties
Biomedical Engineer
Store in Charge
Purchase Officer/ Finance Officer
NS
HEMC should hold a meeting, at least once in a quarter to discuss regarding the equipment functionality and its maintenance.

## Slide 11: Hospital Equipment Management Committee (HEMC)
Functions:
To receive proposal and evaluation with justification for need analysis and cost-benefit analysis
Receipt, installation and commissioning of new equipment
Training of user staff in operation and maintenance
Preparation of equipment master plan with scheduled preventive maintenance and calibration.
Review of inventory, list out functional and non-functional equipment. Coordinating in equipment audit
Mentoring activities of equipment maintenance, complaint management, ensuring preventive maintenance, calibration as per OEM recommendation
Proposal for condemnation
Monitoring the performance of equipment, and break-down maintenance

## Slide 12: Equipment Audit
Equipment audit is periodic evaluation system to measure the quality of performance of the medical equipment.
The focus of the audit is to assess the current status of the medical equipment, carry out equipment census, analyze the records of break down, preventive maintenance, service engineer report, record of feedback from user, performance of equipment, electrical reasons if any etc.
It should analyze the reason for dysfunctional equipment which could be:
Want of minor repair
Lack of PM
Lack of Corrective Maintenance
Lack of Essential Spares
Electrical faults
Unfavorable condition
Mishandling, untrained staff
Purchase without justifiable demand
False reporting, willful damage and overuse than rated

## Slide 13: Equipment Audit
Audit Committee: A committee is constituted under the chairmanship of the hospital administrator, and members representing the user departments. Maintenance history is the main basis of the equipment audit.
The BIS (1994), has described following advantages:
Evaluate the concurrent performance and utilization of equipment
It provides a satisfactory mechanism to assist the process of condemnation
To identify inadequacies and recommend remedial measures
The audit reports provide an objective criteria of procurement of equipment
Cost per reportable result and cost-effectiveness can be evaluated.

## Slide 14: Equipment Audit
Stages involved in equipment audit:
Justification for procurement of equipment
Selection of suppliers
Ensuring availability of the essential spares, after sales service, maintenance service, and training of the users.
Installation and commissioning of equipment

## Slide 15: Common mistakes in hospital for Medical Devices:
Not creating the equipment management committee
Not having an equipment audit program
Delaying the insurance cover/ maintenance contract to save the money
Trying to save the money by postponing the appointment of maintenance staff, especially biomedical engineer.
Buying a costly equipment just for prestige (and without need analysis) and later forcing the doctors to create the demand for that equipment
Not allocating adequate budget for maintenance staff and spare parts
Not following statutory and regulatory framework
Not reporting incidents related to medical devices
9. Not doing calibration of equipment wherein re-calibration is required.

## Slide 16: Indicators for quality of equipment management program
1. Operational Efficiency:
a. Equipment downtime
b. Frequency of breakdown
c. Postponement of procedure
d. Incidence of hazard
e. Frequency of complaints
f. Cost of repairs
2. Utilization indicators:
a. No of tests carried out per month
b. Rate of return from the equipment
c. Equipment utilization quotient

## Slide 17: What is Medical Device?
As per WHO, ‘Medical device‘ means any instrument, apparatus, implement, machine, appliance, implant, reagent for in-vitro use, software, material or other similar or related article, intended by the manufacturer to be used, alone or in combination, for human beings, for one or more of the specific medical purpose(s) of:
1. diagnosis, prevention, monitoring, treatment or alleviation of disease | 2. diagnosis, monitoring, treatment, alleviation of or compensation for an injury
3. investigation, replacement, modification, or support of the anatomy or of a physiological process | 4. supporting or sustaining life
5. control of conception | 6. disinfection of medical devices
7. providing information by means of in-vitro examination of specimens derived from the human body | 8. aids for persons with disabilities
9. devices incorporating animal and/or human tissues | 10. devices for in-vitro fertilization or assisted reproduction technologies
11. disinfection substances; | 12. and does not achieve its primary intended action by pharmacological, immunological or metabolic means, in or on the human body, but which may be assisted in its intended function by such means

## Slide 18: What is Medical Device?
As per MDR, 2017, ‘Medical device’ means –
Substances used for in vitro diagnosis and surgical dressings, surgical bandages, surgical staples, surgical sutures, ligatures, blood and blood component collection bag with or without anticoagulant covered under sub-clause (i) [all medicine for internal or external use of human beings or animals and all substances intended to be used for or in the diagnosis, treatment, mitigation or prevention of any disease or disorder in human beings or animals including preparations applied on human body for the purpose of repelling insects like mosquitoes]

## Slide 19: What is Medical Device?
Substances including mechanical contraceptives (condoms, intrauterine devices, and tubal rings), disinfectants and insecticides notified in the Official Gazette under sub-clause (ii) [such substances (other than food) intended to affect the structure or any function of the human body or intended to be used for the destruction of vermins or insects which cause disease or disorder in human beings or animals, as may be specified from time to time by the central government],
Devices notified from time to time under sub-clause (iv)[such devices intended for internal or external use in the diagnostics, treatment, mitigation or prevention of disease or disorder in human beings or animals, as may be specified y the Central Government], of clause (b) of section 3 of the Act;

## Slide 20: What is Calibration?
Calibration is the process of adjusting and verifying the accuracy of medical devices by comparing their output to a known standard or reference.
According to the ISO, calibration is the set of operations that establish, under specified conditions, the relationship between values indicated by a measuring instrument, a measuring system or values represented by a material measure, and the corresponding known values of a measured.

## Slide 21: Why Calibration is Critical for Patient Safety?
Ensuring Patient Safety:
Calibration ensures that medical devices produce results that meet strict accuracy standards, which in turn protects patients from misdiagnosis or incorrect treatment. Mis-calibrated devices (like thermometers, ECG machines, or blood pressure monitors) can lead to incorrect readings, which may result in misdiagnosis or delayed treatment.
Effective Treatment
Devices such as infusion pumps or ventilators must deliver precise doses or volumes. Even small deviations can be harmful.
Regulatory Compliance
Health authorities (e.g., FDA, WHO, ISO) require regular calibration to maintain certifications and licenses.
Risk Mitigation
Regular calibration reduces the risk of device failure and medical errors, which are major contributors to patient harm.
Quality Control
Testing and calibration helps ensure that medical devices are accurate and effective, and that they last a long time.

## Slide 22: Gaps in the Testing and Calibration of Biomedical Equipment
Limited Technical Expertise among Service Providers
Inadequate Infrastructure
Shortage of Skilled Personnel
Lack of Standardization
Limited Awareness
Disparities in Rural Healthcare
Remote and Challenging Terrains

## Slide 23: Standards in Testing and Calibration of Biomedical Equipment
Standards are established sets of guidelines, criteria, or specifications that act as reference benchmarks to guarantee the accuracy, precision, reliability, and traceability of measurements, hence contributing to safety and quality assurance.
There are two types of standards : written standard and measurement standard.
Written standards are often established by global organizations such as International Organization for Standardization (ISO), International Electrochemical Commission (IEC), Institute of Electrical and Electronics Engineers (IEEE), International Laboratory Accreditation Cooperation (ILAC) and others, providing a universally recognized and accepted framework.
In India, Bureau of Indian Standards (BIS) is an exclusive standardization body for development of Indian Standards (IS), as per the country’s requirement. BIS is a member of the International Organization for Standardization (ISO) and International Electrotechnical Commission (IEC).

## Slide 24: Standards in Testing and Calibration of Biomedical Equipment
BIS has a dedicated division dealing with medical devices under the name “Medical Equipment and Hospital Planning Division (MHD).”
The most widely adopted written quality management standards are IS/ISO 9001:2015 (Quality Management System), IS/ISO/IEC 17025:2017 (General requirements for the competence of testing and calibration laboratories) & IS/ISO/IEC 13485:2016 (Medical Devices- Quality Management Systems, Requirements for regulatory purpose).
Testing & calibration of medical devices should be based on established standards. For example, for testing and calibration of cardiac defibrillator should be based on IEC 60601-1:2012 (Medical electrical equipment: Part 1 General Requirements for basic safety and essential performance) & IEC 60601-2-4: 2010 (Particular requirements for the basic safety and essential performance, Cardiac defibrillators).
If there are no relevant Standard of any medical device that has been laid down by BIS or notified by MoHFW, device shall conform to the standard laid down by the ISO or IEC, or by any other pharmacopeial standards. And in case of the standards which have not been specified by any of the options given above, then the device shall conform to the validated manufacturer’s standards.

## Slide 25: Organizations Contributing to Biomedical Equipment Testing and Calibration
In India, medical device regulations (National Medical Device Policy, 2023 and New Drugs, Medical Devices and Cosmetics Bill, 2022) emphasize on the need for quality control and periodic calibration assessments of medical equipment in order to reduce errors associated with these devices

## Slide 26: CSIR-NPL (Council of Scientific and Industrial Research – National Physical Laboratory, www.nplindia.org )
CSIR-NPL is mandated to be India’s “National Metrology Institute” by the act of Parliament (The Gazette of India) and is custodian of “National Standards of Measurements.” CSIR-NPL is responsible for the realization, establishment, up-gradation, maintenance and dissemination of standards at par to international level, providing metrological traceability across the country
The CSIR-NPL has following facilities for testing/calibration of medical device, medical device simulator/analyser : Defibrillator & Defibrillator Analyser calibration, Infusion Device & Infusion Device Analyser, Non-Invasive Blood Pressure (NIBP) devices calibration/testing system, Clinical Thermometer testing/calibration facility, ECG Simulator calibration facility.

## Slide 27: National Accreditation Board for Testing and Calibration Laboratories (NABL)
NABL is an accreditation body, with its accreditation system established in accordance with ISO/ IEC 17011. It undertakes the assessment and accreditation of testing (including medical), calibration, proficiency testing provider and reference material producer in India and abroad.
NABL offers voluntary accreditation services to calibration laboratories in accordance with ISO/IEC 17025. The scope of NABL accreditation services for calibration laboratory encompass multiple disciplines, including Mechanical, Electro-technical, Fluid Flow, Thermal, Optical, Radiological and Medical Devices.
NABL provide accreditation to calibration laboratories that adheres to stringent quality and calibration standards and procedure.
NABL provide accreditation to a large number of laboratories while ensuring metrological traceability to SI units.

## Slide 28: Standardization of Testing and Quality Certification (STQC) Directorate
Standardization Testing and Quality Certification (STQC) Directorate is an attached office of the Ministry of Electronics and Information Technology, Government of India, provides quality assurance services in the area of Electronics and IT through countrywide network of laboratories and centers.
STQC offers calibration services to industry and organizations in the domains: Electro Technical Calibration, Non-Electrical Calibration, High Precision Calibration, Onsite Calibration, and Medical Equipment Calibration.
There are 4 regional laboratories, 10 state level laboratories and 2 high precision calibration centres at Delhi and Bangaluru. The STQC laboratories are equipped with the state-of-the-art calibrating facilities.
Medical equipment calibration services are provided by STQC laboratories at ERTL (North) Delhi and ERTL (South) Thiruvananthapuram. https://www.stqc.gov.in/labs-centres

## Slide 29: Documentation of Calibration Results
Information obtained through calibration can be conveyed through a statement, calibration function, calibration diagram, calibration curve, or a calibration table. In some cases, it may consist of an additive or multiplicative correction of the indication with associated measurement uncertainty.
Calibration Diagram: graphical expression of the relation between indication and corresponding measurement result. This concept pertains to a calibration when the instrumental measurement uncertainty is large in comparison with the measurement uncertainties associated with the quantity values of measurement standards (IEC/ISO Guide99, Definition 4.30).
Calibration Curve: expression of the relation between indication and corresponding measured quantity value. A calibration curve expresses a one-to-one relation that does not supply a measurement result as it bears no information about the measurement uncertainty (IEC/ISO Guide99, Definition 4.31).

## Slide 30: Calibration Interval
A calibration interval is the predefined period between two consecutive scheduled calibrations. The primary purpose is to ensure that the instrument continues to provide accurate and reliable measurements, in accordance with the reference standards or manufacturer’s specifications.
The frequency of calibration may vary depending on factors such as instrument type, manufacturer recommendations, instrument usage, criticality of use, environmental conditions, required accuracy and precision, maximum permissible errors, stability of the equipment and regulatory requirements etc.
It is recommended that healthcare facilities consult with equipment manufacturers, regulatory bodies, and local experts to establish appropriate calibration schedules for their specific biomedical equipment and circumstances.
Biomedical Equipment at various healthcare facilities (as per IPHS 2022) has been classified according to the calibration requirements, relevant standards, and the reference equipment necessary for calibration along with their suggestive calibration schedule are annexed.

## Slide 31: Equipment maintenance
BIS (1994) has defined the equipment maintenance as “the combination of all technical and administrative actions, including supervision, actions, intended to retain an item in, or restore it, to a state in which it can perform a required function”.
Maintenance is an action or combination of actions carried out to retain equipment in or restore it to an acceptable condition.
The maintenance program in fact starts from the planning of the equipment and carries value till the equipment is disposed.

## Slide 32: Maintenance Policy
Each hospital should develop the maintenance policy
The basic concept of the maintenance policy is to support the life and health of patient while keeping the reliability and safety of the equipment

## Slide 33: Objectives of maintenance program
Uninterrupted operation or services in a hospital
Extend the useful life of assets
Assure the optimum availability of installed plant & equipment for production (service) and to obtain the maximum possible return on investment of these items.
Ensure the operational readiness of all equipment for emergency use at all times.
Ensure the safety of patients, person using & operating the equipment
Provide a systematic approach to work in the department.
Provide a sound method of ascertaining the validity of budgetary request.

## Slide 34: Types of maintenance program
Preventive maintenance: At predetermined interval/ manufacturer’s recommendation. Good quality spares/accessories s essential. Advantages: reduction in down-time ensures safety, extend life and availability during emergency
Breakdown/ Corrective maintenance: Carried out after fault recognition. To analyze the reason for break-down/ shut down and ensure early rectification. The user should immediately contact the service engineer/ service provider and log his complaint using Toll Free number provided by service provider/ OEM for call resolution. After repair they should provide a guarantee/ warranty certificate stating that the equipment serviced by them will render trouble free service for a defined period. In case the break-down is in warranty period, it should be repaired free of cost.
Predictive maintenance
Total productive maintenance
Other system of maintenance (Planned, running, shutdown, overhauling)

## Slide 35: Contractual Maintenance System
Advantages:
There is no capital investment in equipment, tool etc.
Well trained specialists are hired for each job.
There are no in house personnel problems such as strike, etc.
Disadvantages:
It is no longer possible to control when and how each job will be completed
Costs are higher because the contracting firms must make a profit
They charge emergency cost for repair and in large countries like India they can not have an economical servicing network

## Slide 36: Equipment Management Information System
The information system should preferably be computerized
The data can be identified on the basis of the history sheet of the equipment
This data sheet can be maintained manually also in small settings, but if maintained on a computer database like access, fox pro or readily available software of the inventory control, the results are really rewarding.
BIS has recommended following contents in the history-sheet of equipment which should be made available at all the time with owner:
1. Name of the equipment | 2. Date of purchase | 3. Cost | 4. Supplier details | 5. Manufacturer details
6. Date of installation | 7. Location | 8. Environmental Control | 9. Spare parts | 10. Manual
11. After sale service | 12. Warranty | 13. Life of equipment | 14. Depreciation | 15. Charges of tests
16. Use co-efficient | 17. Down time/ Up time | 18. Cost of maintenance | 19. Date of condemnation | 20. Date of replacement

## Slide 37: Decommissioning and disposal
Equipment condemnation and disposal
Assessment of life of the equipment can be done on the following basis: | 
1. Built up structure & technology | 5. Handling of the equipment
2. Number of years | 6. Change in the technology
3. Maintenance & repair | 7. Obsolescence of technology
4. Availability of spare parts | 
Minimum criteria as recommended by BIS for condemnation and disposal | 
1. Non-functional & beyond economical repair | 4. Functional but hazardous
2. Non-functional & obsolete | 5. Functional but no longer required
3. Functional but obsolete | 
Disposal | 
1. By auction | 3. Transfer to other hospital
2. Buy back to firm | 4. Scrap

## Slide 38: Thank You